## Raw Data
Samples | Coated PMP fibres (LN021/117/2) | [11 empty] | Coated PMP fibres (LN021/117/2) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740093 | Group | S/N | ref | ANOVA | T-Test | PC Coated | EtO | Temp Cycle | 2 Yrs Simulated Age | 14 Day Sim Use
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | SU | 1 | ECMO (0740099) | LN021/117/3 |  |  | X
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B |  |  | 1 | ECMO (0740127) | LN021/117/7 |  |  | X
Image 1 | 2378.4816 | 2254.315 | 3265.7236 | 2562.7405 | 2413.026 | 1696.616 | 2178.5054 | 2424.2884 | 2418.9599 | 1598.537 | 2852.2896 | 1935.9269 | 1892.7323 | 2334.7448 | 2053.7296 | 2390.2251 | 2912.6675 | 2417.3037 | 3258.5025 | 2989.9897 | 641.1342 | 614.098 |  |  | 2 | ECMO (0740093) | LN021/117/2 |  |  | X |  |  |  | X
Image 2 | 2341.7996 | 1627.5081 | 3025.1484 | 2774.7128 | 2315.0452 | 1797.7282 | 1678.4096 | 1725.3871 | 2038.941 | 2435.637 | 2694.0107 | 2171.5758 | 2643.1281 | 2685.5554 | 2664.7203 | 2307.5829 | 2614.8823 | 3263.8718 | 3146.5223 | 2788.713 | 683.2529 | 542.851 |  |  | 2 | ECMO (0740128) | LN021/117/8 |  |  | X |  |  |  | X
Image 3 | 1869.9868 | 2227.8814 | 2907.7031 | 1222.9905 | 1968.8736 | 1333.5361 | 1601.4655 | 1910.9667 | 1786.8704 | 1822.8851 | 3045.5644 | 2920.9688 | 2183.2422 | 1629.778 | 2445.6591 | 1964.8862 | 3180.8275 | 2865.0431 | 3319.7046 | 3168.7609 | 780.1471 | 739.8454 |  |  | 3 | ECMO (0740124) | LN021/117/6 |  |  | X | X | X
Image 4 | 1260.284 | 1786.6794 | 3227.0805 | 930.9537 | 2073.3394 | 1252.4967 | 1402.8272 | 2219.6376 | 2592.8028 | 2050.4665 | 2587.9352 | 1892.7111 | 2324.2728 | 1814.142 | 2776.003 | 2172.4622 | 2792.6245 | 2760.3326 | 2704.0868 | 3152.3975 | 633.1879 | 854.5569 |  |  | 3 | ECMO (0740129) | LN021/117/9 |  |  | X | X | X
average | 1962.638 | 1974.096 | 3106.4139 | 1872.8494 | 2192.5711 | 1520.0943 | 1715.3019 | 2070.07 | 2209.3935 | 1976.8814 | 2794.95 | 2230.2956 | 2260.8438 | 2116.055 | 2485.028 | 2208.7891 | 2875.2504 | 2826.6378 | 3107.2041 | 3024.9653 | 684.4306 | 687.8378 |  |  | 4 | ECMO (0740104) | LN021/117/4 |  |  | X | X | X |  | X
SD | 522.3584 | 315.2659 | 169.342 | 930.7329 | 206.4437 | 267.4902 | 329.9079 | 311.9622 | 364.4633 | 357.1791 | 199.2814 | 476.4779 | 311.8186 | 482.9675 | 318.6032 | 185.7306 | 237.6243 | 348.616 | 278.1469 | 176.9695 | 67.4869 | 137.7906 |  |  | 4 | ECMO (0740106) | LN021/117/5 |  |  | X | X | X |  | X
CV | 0.26615 | 0.1597 | 0.054514 | 0.49696 | 0.094156 | 0.17597 | 0.19233 | 0.1507 | 0.16496 | 0.18068 | 0.071301 | 0.21364 | 0.13792 | 0.22824 | 0.12821 | 0.084087 | 0.082645 | 0.12333 | 0.089517 | 0.058503 | 0.098603 | 0.20032
F-test | 0.001699 |  | 7.2596e-06 |  | 0.0011954 |  | 0.0037192 |  | 0.0034766 |  | 0.00075467 |  | 0.0021736 |  | 0.013922 |  | 0.01562 |  | 0.055239
T-test | 2.3891e-05 |  | 0.00074602 |  | 7.0674e-05 |  | 1.299e-05 |  | 4.4765e-06 |  | 5.1955e-06 |  | 5.3488e-06 |  | 1.0876e-07 |  | 8.8019e-09 |  | 1.1878e-13
Samples | Coated PET fibres (LN021/117/2) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/2 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2637.8904 | 2988.8879 | 2836.8024 | 2674.5104 | 2856.865 | 2644.6885 | 699.6462 | 747.8233 | 3381.4997 | 947.8812 | 644.4613 | 639.8262
Image 2 | 2432.2594 | 2239.226 | 2155.7951 | 2756.8942 | 2784.8678 | 3181.5742 | 647.8825 | 857.5156 | 941.6554 | 1014.378 | 727.7484 | 625.3321
Image 3 | 2881.9553 | 2685.3145 | 2449.3474 | 2951.33 | 3149.4272 | 2988.1893 | 669.1921 | 724.5108 | 1393.53 | 933.8401 | 653.6579 | 664.9418
Image 4 | 2701.0733 | 2167.6967 | 2313.1986 | 2779.2422 | 2355.0447 | 1632.1424 | 781.5203 | 834.7282 | 1523.5612 | 1232.8966 | 786.2552 | 515.9295
average | 2663.2946 | 2520.2813 | 2438.7859 | 2790.4942 | 2786.5512 | 2611.6486 | 699.5603 | 791.1445 | 1810.0616 | 1032.249 | 703.0307 | 611.5074
SD | 185.5284 | 387.3569 | 291.1956 | 116.2984 | 328.0364 | 689.7154 | 58.6239 | 64.8625 | 1076.898 | 138.3005 | 66.8464 | 65.7862
CV | 0.069661 | 0.1537 | 0.1194 | 0.041677 | 0.11772 | 0.26409 | 0.083801 | 0.081986 | 0.59495 | 0.13398 | 0.095083 | 0.10758
F-test | 0.0020218 |  | 0.002696 |  | 5.1658e-05 |  |  |  | 2.5386e-06
T-test | 1.3645e-07 |  | 7.8324e-08 |  | 9.7395e-06 |  |  |  | 0.034266
Samples | Coated PMP fibres (LN021/117/3) | [11 empty] | Coated PMP fibres (LN021/117/3) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740099
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | None
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 3181.6689 | 3241.5791 | 2714.9583 | 3082.3056 | 3027.6035 | 2797.2115 | 2862.8529 | 2953.9138 | 2718.2051 | 2833.0384 | 2981.0669 | 2905.1643 | 3024.5048 | 3156.2699 | 2814.4034 | 3062.9868 | 2584.1546 | 2930.2636 | 3196.3815 | 3365.4069 | 641.1342 | 614.098
Image 2 | 2421.6522 | 3266.4638 | 2345.1136 | 2864.8576 | 2617.5795 | 2784.8612 | 3091.1175 | 2951.7409 | 2691.9955 | 2756.0615 | 2939.9918 | 2805.9342 | 3012.0608 | 3194.3966 | 2950.5432 | 3010.3628 | 2722.9034 | 2459.8034 | 3116.0255 | 3240.9634 | 683.2529 | 542.851
Image 3 | 2719.2064 | 2884.1641 | 2135.8904 | 2987.797 | 2664.6041 | 2639.5694 | 3065.8495 | 3183.5751 | 2865.2405 | 2931.8701 | 2915.9741 | 2988.0708 | 3065.9233 | 3105.7137 | 2777.2282 | 3082.221 | 2665.7225 | 2895.6955 | 3020.9317 | 3042.5062 | 780.1471 | 739.8454
Image 4 | 2756.4959 | 3072.0327 | 2601.8319 | 3069.8455 | 2550.5601 | 2855.9281 | 2890.4061 | 3238.4136 | 2762.9864 | 2920.9567 | 3001.3781 | 2608.1564 | 3175.8418 | 2964.8379 | 3037.4138 | 2770.8669 | 2454.633 | 2697.205 | 3007.9828 | 3311.8118 | 633.1879 | 854.5569
average | 2769.7558 | 3116.0599 | 2449.4486 | 3001.2014 | 2715.0868 | 2769.3925 | 2977.5565 | 3081.9108 | 2759.6069 | 2860.4817 | 2959.6027 | 2826.8314 | 3069.5827 | 3105.3045 | 2894.8972 | 2981.6094 | 2606.8534 | 2745.7419 | 3085.3304 | 3240.1721 | 684.4306 | 687.8378
SD | 312.8256 | 177.0974 | 260.0773 | 100.0986 | 213.5352 | 91.9341 | 117.5356 | 150.7273 | 76.2787 | 82.4828 | 38.7038 | 163.6965 | 74.4875 | 100.4422 | 120.7376 | 143.7406 | 116.3607 | 216.5263 | 88.3259 | 141.2898 | 67.4869 | 137.7906
CV | 0.11294 | 0.056834 | 0.10618 | 0.033353 | 0.078648 | 0.033196 | 0.039474 | 0.048907 | 0.027641 | 0.028835 | 0.013077 | 0.057908 | 0.024266 | 0.032345 | 0.041707 | 0.048209 | 0.044636 | 0.078859 | 0.028628 | 0.043606 | 0.098603 | 0.20032
F-test | 0.010009 |  | 0.0041199 |  | 0.27559 |  | 0.43184 |  | 0.80504 |  | 0.50025 |  | 0.64991 |  | 0.49619 |  | 0.15726 |  | 0.43259
T-test | 1.6053e-08 |  | 1.8984e-07 |  | 2.1329e-14 |  | 1.0967e-15 |  | 1.8941e-16 |  | 1.6955e-15 |  | 2.1021e-17 |  | 1.31e-15 |  | 1.2956e-13 |  | 5.0715e-16
Samples | Coated PET fibres (LN021/117/3) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/3 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2983.6599 | 2793.948 | 2694.5615 | 3060.012 | 2856.8003 | 2501.6353 | 699.6462 | 747.8233 | 740.5048 | 1031.7137 | 644.4613 | 639.8262
Image 2 | 2658.9663 | 2851.2659 | 3177.358 | 3065.4306 | 2591.8847 | 2205.6745 | 647.8825 | 857.5156 | 1109.3175 | 1931.6676 | 727.7484 | 625.3321
Image 3 | 2536.6966 | 2990.1245 | 3058.9713 | 2932.5879 | 2525.7359 | 2115.3129 | 669.1921 | 724.5108 | 2792.2531 | 1281.8028 | 653.6579 | 664.9418
Image 4 | 2864.7842 | 2728.8824 | 3176.8622 | 3019.5863 | 2657.1741 | 2122.2132 | 781.5203 | 834.7282 | 637.3471 | 1624.8382 | 786.2552 | 515.9295
average | 2761.0267 | 2841.0552 | 3026.9382 | 3019.4042 | 2657.8987 | 2236.209 | 699.5603 | 791.1445 | 1319.8556 | 1467.5056 | 703.0307 | 611.5074
SD | 200.8909 | 111.2471 | 228.4759 | 61.3855 | 143.0469 | 181.6538 | 58.6239 | 64.8625 | 1002.2884 | 393.531 | 66.8464 | 65.7862
CV | 0.072759 | 0.039157 | 0.075481 | 0.02033 | 0.05382 | 0.081233 | 0.083801 | 0.081986 | 0.75939 | 0.26816 | 0.095083 | 0.10758
F-test | 0.073017 |  | 0.076292 |  | 0.0031493 |  |  |  | 7.0921e-06
T-test | 9.0337e-15 |  | 1.9748e-15 |  | 1.2708e-07 |  |  |  | 0.021788
Samples | Coated PMP fibres (LN021/117/4) | [11 empty] | Coated PMP fibres (LN021/117/4) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740104
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | ETO,TC,SU
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 2863.0418 | 2830.9147 | 2455.8632 | 2360.545 | 3249.1164 | 3596.5865 | 3156.4845 | 2556.4139 | 2690.906 | 2632.5657 | 3029.7786 | 2492.6009 | 2283.7682 | 2267.1574 | 2687.9215 | 2621.7847 | 2411.5831 | 2094.424 | 2669.9396 | 2730.4412 | 641.1342 | 614.098
Image 2 | 2577.2988 | 2460.8748 | 2064.2471 | 2073.2207 | 2824.4594 | 3395.4058 | 3100.7939 | 2972.4185 | 2046.2038 | 2983.9751 | 2647.4447 | 2684.8112 | 2194.9513 | 2218.0937 | 2934.9545 | 2640.4016 | 2555.6143 | 2382.0022 | 2524.6578 | 2476.9588 | 683.2529 | 542.851
Image 3 | 2555.8466 | 2755.6042 | 2502.4063 | 1938.4847 | 3617.3592 | 3047.6747 | 3151.2483 | 2697.8175 | 2672.8452 | 2285.0791 | 2672.7286 | 2602.4578 | 2624.5492 | 2625.8965 | 2668.653 | 2788.2804 | 2666.1315 | 2441.9894 | 2380.8934 | 2388.4198 | 780.1471 | 739.8454
Image 4 | 2879.7116 | 2770.6969 | 2638.4445 | 2761.4569 | 3496.3627 | 2741.602 | 3289.3558 | 2801.1216 | 2359.0991 | 2304.1073 | 2740.1837 | 2759.5334 | 2501.2015 | 2461.5132 | 2843.9224 | 2916.4783 | 2213.8255 | 2595.8617 | 2266.2517 | 2413.9156 | 633.1879 | 854.5569
average | 2718.9747 | 2704.5227 | 2415.2403 | 2283.4268 | 3296.8244 | 3195.3173 | 3174.4706 | 2756.9429 | 2442.2635 | 2551.4318 | 2772.5339 | 2634.8508 | 2401.1176 | 2393.1652 | 2783.8628 | 2741.7362 | 2461.7886 | 2378.5693 | 2460.4356 | 2502.4339 | 684.4306 | 687.8378
SD | 176.3278 | 165.6578 | 246.4849 | 364.0658 | 350.2203 | 378.0248 | 80.6011 | 175.2039 | 304.8336 | 329.5399 | 175.907 | 114.4933 | 196.8196 | 187.4062 | 127.6896 | 138.273 | 195.4189 | 209.7526 | 175.1673 | 156.4939 | 67.4869 | 137.7906
CV | 0.064851 | 0.061252 | 0.10205 | 0.15944 | 0.10623 | 0.11831 | 0.02539 | 0.06355 | 0.12482 | 0.12916 | 0.063446 | 0.043453 | 0.08197 | 0.078309 | 0.045868 | 0.050433 | 0.079381 | 0.088184 | 0.071194 | 0.062537 | 0.098603 | 0.20032
F-test | 0.25145 |  | 0.01064 |  | 0.0045128 |  | 0.0244 |  | 0.0099679 |  | 0.2692 |  | 0.15435 |  | 0.57477 |  | 0.10664 |  | 0.2724
T-test | 3.2897e-14 |  | 1.8218e-07 |  | 2.6267e-08 |  | 1.9183e-09 |  | 1.0193e-07 |  | 2.9365e-14 |  | 1.077e-12 |  | 2.6835e-15 |  | 2.0976e-12 |  | 1.4253e-13
Samples | Coated PET fibres (LN021/117/4) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/4 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2358.5202 | 2459.4463 | 1690.7528 | 1650.5145 | 1794.1319 | 2101.3408 | 699.6462 | 747.8233 | 825.4107 | 3326.967 | 644.4613 | 639.8262
Image 2 | 2218.0354 | 2260.195 | 1936.2709 | 1870.7535 | 1804.1229 | 1878.2711 | 647.8825 | 857.5156 | 1254.1573 | 3033.3958 | 727.7484 | 625.3321
Image 3 | 2279.7841 | 2692.3159 | 1700.6422 | 1843.2641 | 1972.68 | 2030.2514 | 669.1921 | 724.5108 | 656.1189 | 3844.7955 | 653.6579 | 664.9418
Image 4 | 2505.5383 | 2347.4931 | 1842.4543 | 1873.9399 | 1255.8239 | 2028.731 | 781.5203 | 834.7282 | 1393.2361 | 2726.0069 | 786.2552 | 515.9295
average | 2340.4695 | 2439.8626 | 1792.53 | 1809.618 | 1706.6897 | 2009.6486 | 699.5603 | 791.1445 | 1032.2308 | 3232.7913 | 703.0307 | 611.5074
SD | 124.159 | 187.0194 | 118.2594 | 106.9593 | 311.5393 | 93.908 | 58.6239 | 64.8625 | 348.2389 | 476.0978 | 66.8464 | 65.7862
CV | 0.053049 | 0.076652 | 0.065973 | 0.059106 | 0.18254 | 0.046729 | 0.083801 | 0.081986 | 0.33737 | 0.14727 | 0.095083 | 0.10758
F-test | 0.073097 |  | 0.40318 |  | 0.0034482 |  |  |  | 1.5027e-07
T-test | 1.954e-13 |  | 1.4774e-12 |  | 3.0097e-06 |  |  |  | 0.011882
Samples | Coated PMP fibres (LN021/117/5) | [11 empty] | Coated PMP fibres (LN021/117/5) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740106
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | ETO,TC,SU
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 2464.3736 | 2673.0178 | 2730.4269 | 3054.03 | 2991.512 | 2992.4466 | 2922.9644 | 3022.9661 | 2901.0098 | 3268.9635 | 2917.7791 | 3150.8024 | 2851.1066 | 2874.0797 | 2933.5803 | 2703.5108 | 3055.6788 | 2750.366 | 2582.6848 | 2700.829 | 641.1342 | 614.098
Image 2 | 2626.9398 | 2655.3278 | 2434.7131 | 3296.4884 | 2874.1021 | 2897.7052 | 3051.0151 | 3046.8018 | 3473.8135 | 3078.1916 | 2844.9032 | 2989.646 | 3194.9068 | 2895.4717 | 2802.2519 | 2662.3013 | 2906.6396 | 2989.5273 | 2654.1138 | 2943.9209 | 683.2529 | 542.851
Image 3 | 2422.9786 | 2907.9757 | 2790.1312 | 2728.0367 | 3202.4331 | 3257.8077 | 2575.0968 | 3045.981 | 2784.5701 | 3334.2308 | 2730.3791 | 2772.7309 | 3018.3914 | 2920.2968 | 3044.1549 | 3002.0897 | 2627.0618 | 3173.2125 | 2763.1868 | 2779.6142 | 780.1471 | 739.8454
Image 4 | 2817.43 | 3110.6741 | 2449.2359 | 2954.0499 | 2596.2912 | 3237.9554 | 2431.5232 | 2828.9618 | 3306.8878 | 3579.0621 | 2706.9395 | 2856.1208 | 2472.2987 | 2582.9279 | 3056.3593 | 3051.6681 | 2805.1343 | 2665.7889 | 2976.891 | 2011.8561 | 633.1879 | 854.5569
average | 2582.9305 | 2836.7488 | 2601.1268 | 3008.1513 | 2916.0846 | 3096.4787 | 2745.1499 | 2986.1777 | 3116.5703 | 3315.112 | 2800.0002 | 2942.325 | 2884.1759 | 2818.194 | 2959.0866 | 2854.8925 | 2848.6286 | 2894.7237 | 2744.2191 | 2609.055 | 684.4306 | 687.8378
SD | 179.4131 | 215.8933 | 185.4773 | 235.6794 | 252.7944 | 179.2357 | 290.0869 | 105.3912 | 326.871 | 206.7979 | 98.9874 | 165.2225 | 308.385 | 157.977 | 118.2459 | 200.3293 | 180.0173 | 230.7914 | 171.9576 | 410.81 | 67.4869 | 137.7906
CV | 0.069461 | 0.076106 | 0.071307 | 0.078347 | 0.08669 | 0.057884 | 0.10567 | 0.035293 | 0.10488 | 0.06238 | 0.035353 | 0.056154 | 0.10692 | 0.056056 | 0.03996 | 0.070171 | 0.063194 | 0.079728 | 0.062662 | 0.15746 | 0.098603 | 0.20032
F-test | 0.045678 |  | 0.011356 |  | 0.049886 |  | 0.035378 |  | 0.016578 |  | 0.33434 |  | 0.0445 |  | 0.22973 |  | 0.10582 |  | 0.0098218
T-test | 1.0273e-09 |  | 2.1071e-08 |  | 2.1389e-10 |  | 1.0691e-09 |  | 1.9822e-09 |  | 5.6483e-15 |  | 5.8648e-10 |  | 1.151e-14 |  | 8.9313e-14 |  | 4.7697e-08
Samples | Coated PET fibres (LN021/117/5) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/5 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2141.8304 | 2337.538 | 2084.0527 | 1814.2796 | 1494.5445 | 1560.447 | 699.6462 | 747.8233 | 641.2336 | 592.0228 | 644.4613 | 639.8262
Image 2 | 2152.7657 | 1987.6483 | 1985.3927 | 2843.7943 | 2072.7741 | 2165.2823 | 647.8825 | 857.5156 | 1232.1345 | 2497.8124 | 727.7484 | 625.3321
Image 3 | 1418.2498 | 1391.5876 | 2170.8926 | 2441.7669 | 2003.393 | 1768.1913 | 669.1921 | 724.5108 | 637.6588 | 3253.4964 | 653.6579 | 664.9418
Image 4 | 1749.6397 | 1691.8753 | 2416.7401 | 2331.1295 | 2213.7531 | 1955.6123 | 781.5203 | 834.7282 | 830.5239 | 721.9056 | 786.2552 | 515.9295
average | 1865.6214 | 1852.1623 | 2164.2695 | 2357.7426 | 1946.1162 | 1862.3832 | 699.5603 | 791.1445 | 835.3877 | 1766.3093 | 703.0307 | 611.5074
SD | 352.2955 | 404.8732 | 184.5868 | 424.0152 | 313.512 | 258.5067 | 58.6239 | 64.8625 | 279.4186 | 1318.6546 | 66.8464 | 65.7862
CV | 0.18884 | 0.21859 | 0.085288 | 0.17984 | 0.1611 | 0.1388 | 0.083801 | 0.081986 | 0.33448 | 0.74656 | 0.095083 | 0.10758
F-test | 0.0006062 |  | 0.0011139 |  | 0.0032792 |  |  |  | 6.0513e-07
T-test | 2.9914e-05 |  | 1.4658e-06 |  | 2.3767e-06 |  |  |  | 0.11582
Samples | Coated PMP fibres (LN021/117/6) | [11 empty] | Coated PMP fibres (LN021/117/6) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740124
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | ETO,TC
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 3054.8287 | 3018.0271 | 1994.7481 | 1549.3609 | 2585.9633 | 2412.2903 | 2452.8677 | 2227.6063 | 2454.5619 | 2460.9795 | 2434.949 | 2596.904 | 2793.2361 | 2385.1557 | 2260.0688 | 2162.61 | 2301.6264 | 2673.0414 | 2884.6605 | 2814.6566 | 641.1342 | 614.098
Image 2 | 3248.4441 | 3167.7966 | 2020.6852 | 1359.0485 | 2534.0187 | 2638.4543 | 2368.374 | 2565.0841 | 2600.5349 | 2295.119 | 2661.1026 | 2442.967 | 2517.3241 | 2359.7423 | 2412.3186 | 2343.4109 | 2405.7324 | 2664.0178 | 2886.5893 | 2750.7563 | 683.2529 | 542.851
Image 3 | 3433.0618 | 3292.5549 | 1811.8248 | 1761.6694 | 2615.4848 | 2502.2486 | 2000.9901 | 2443.4375 | 2526.2762 | 2187.7822 | 2347.7143 | 2546.3872 | 2618.0972 | 2639.5829 | 2554.0428 | 2644.0451 | 2395.4715 | 2458.9507 | 2966.1649 | 1910.2393 | 780.1471 | 739.8454
Image 4 | 3474.1168 | 3130.4837 | 1763.2551 | 2126.3213 | 2525.7994 | 2319.9066 | 2540.621 | 2359.0251 | 2662.8696 | 1971.0051 | 2091.0853 | 2585.4919 | 2274.9572 | 2539.2658 | 2467.8757 | 2341.0929 | 2766.8134 | 2763.9027 | 2602.6896 | 2103.8086 | 633.1879 | 854.5569
average | 3302.6128 | 3152.2156 | 1897.6283 | 1699.1 | 2565.3165 | 2468.225 | 2340.7132 | 2398.7883 | 2561.0606 | 2228.7214 | 2383.7128 | 2542.9375 | 2550.9037 | 2480.9367 | 2423.5765 | 2372.7897 | 2467.4109 | 2639.9781 | 2835.0261 | 2394.8652 | 684.4306 | 687.8378
SD | 192.1473 | 113.1617 | 129.0912 | 328.8818 | 42.7563 | 135.7236 | 237.1488 | 142.0472 | 90.3239 | 205.3016 | 235.5822 | 70.0696 | 216.4209 | 132.2031 | 123.6196 | 199.6854 | 205.0251 | 128.8401 | 159.4783 | 455.5069 | 67.4869 | 137.7906
CV | 0.05818 | 0.035899 | 0.068028 | 0.19356 | 0.016667 | 0.054988 | 0.10131 | 0.059216 | 0.035268 | 0.092116 | 0.09883 | 0.027555 | 0.084841 | 0.053288 | 0.051007 | 0.084156 | 0.083093 | 0.048803 | 0.056253 | 0.1902 | 0.098603 | 0.20032
F-test | 0.20507 |  | 0.025451 |  | 0.87894 |  | 0.13408 |  | 0.043576 |  | 0.13948 |  | 0.18763 |  | 0.26747 |  | 0.13477 |  | 0.0018565
T-test | 2.3459e-15 |  | 9.7759e-07 |  | 4.3264e-15 |  | 1.8606e-12 |  | 5.7626e-09 |  | 8.1062e-13 |  | 2.7096e-13 |  | 2.864e-13 |  | 4.4477e-13 |  | 1.0146e-06
Samples | Coated PET fibres (LN021/117/6) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/6 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 3032.7795 | 3378.5007 | 2421.844 | 2599.4867 | 2655.3467 | 2256.3214 | 699.6462 | 747.8233 | 2942.1534 | 1088.4144 | 644.4613 | 639.8262
Image 2 | 2758.9397 | 2918.6306 | 2217.541 | 2557.0076 | 2366.7832 | 2587.1894 | 647.8825 | 857.5156 | 2318.6672 | 1350.5113 | 727.7484 | 625.3321
Image 3 | 3065.8432 | 3179.9934 | 2047.6246 | 2107.2224 | 2255.9828 | 2659.6132 | 669.1921 | 724.5108 | 2221.5244 | 941.9382 | 653.6579 | 664.9418
Image 4 | 2888.7699 | 2685.4395 | 1602.2146 | 2466.6744 | 2759.1901 | 2098.1068 | 781.5203 | 834.7282 | 2289.4267 | 1541.3655 | 786.2552 | 515.9295
average | 2936.5831 | 3040.6411 | 2072.306 | 2432.5978 | 2509.3257 | 2400.3077 | 699.5603 | 791.1445 | 2442.9429 | 1230.5573 | 703.0307 | 611.5074
SD | 141.1916 | 302.5574 | 348.743 | 223.875 | 236.823 | 267.2212 | 58.6239 | 64.8625 | 335.2854 | 267.3925 | 66.8464 | 65.7862
CV | 0.04808 | 0.099504 | 0.16829 | 0.092031 | 0.094377 | 0.11133 | 0.083801 | 0.081986 | 0.13725 | 0.21729 | 0.095083 | 0.10758
F-test | 0.0097307 |  | 0.00086509 |  | 0.0065611 |  |  |  | 7.3067e-06
T-test | 1.5821e-09 |  | 2.1515e-06 |  | 3.3772e-08 |  |  |  | 0.0020816
Samples | Coated PMP fibres (LN021/117/7) | [11 empty] | Coated PMP fibres (LN021/117/7) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740127
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | None
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 3178.0189 | 2789.5636 | 2038.7641 | 3179.4212 | 2585.2993 | 2970.8303 | 3231.4217 | 2940.9787 | 2504.6697 | 2995.7879 | 2682.9125 | 3102.8217 | 3041.7587 | 3124.1121 | 2670.831 | 3189.7532 | 3051.9136 | 3232.9146 | 3056.0392 | 3088.9803 | 641.1342 | 614.098
Image 2 | 3215.1847 | 2802.5534 | 1844.5947 | 2986.5731 | 2207.5946 | 3070.7198 | 2266.4358 | 3138.6758 | 3097.7052 | 3013.935 | 3024.9971 | 2615.0374 | 2418.1769 | 3294.0299 | 3183.9979 | 3321.2814 | 3056.4814 | 3012.2855 | 3104.331 | 3109.4699 | 683.2529 | 542.851
Image 3 | 3058.777 | 2941.0622 | 2818.1572 | 3276.0139 | 2558.4554 | 2675.6434 | 3026.0793 | 2808.0615 | 3079.7169 | 3056.7054 | 2526.8007 | 2757.7102 | 3241.1228 | 2890.0426 | 3216.8255 | 3015.3547 | 3092.4047 | 3023.6859 | 3198.0748 | 3349.7138 | 780.1471 | 739.8454
Image 4 | 3341.3057 | 2032.6742 | 2771.6182 | 3155.3999 | 2330.7945 | 2892.3471 | 2824.0362 | 2786.1316 | 2815.4724 | 2918.3119 | 3206.8353 | 2587.9224 | 3294.6769 | 2926.4789 | 3330.7027 | 3140.9453 | 3141.0377 | 3302.523 | 3126.7099 | 3313.8626 | 633.1879 | 854.5569
average | 3198.3216 | 2641.4634 | 2368.2835 | 3149.352 | 2420.5359 | 2902.3851 | 2836.9933 | 2918.4619 | 2874.3911 | 2996.185 | 2860.3864 | 2765.8729 | 2998.9338 | 3058.6659 | 3100.5893 | 3166.8336 | 3085.4594 | 3142.8522 | 3121.2887 | 3215.5067 | 684.4306 | 687.8378
SD | 116.3525 | 411.6096 | 499.2989 | 120.3896 | 182.1778 | 167.8629 | 415.143 | 161.9678 | 278.2045 | 57.8566 | 310.859 | 236.6565 | 402.1741 | 187.6047 | 293.3214 | 126.4852 | 41.2401 | 147.0309 | 59.0778 | 135.3245 | 67.4869 | 137.7906
CV | 0.036379 | 0.15583 | 0.21083 | 0.038227 | 0.075263 | 0.057836 | 0.14633 | 0.055498 | 0.096787 | 0.01931 | 0.10868 | 0.085563 | 0.13411 | 0.061335 | 0.094602 | 0.039941 | 0.013366 | 0.046783 | 0.018927 | 0.042085 | 0.098603 | 0.20032
F-test | 0.0014696 |  | 0.00025069 |  | 0.0090815 |  | 0.010933 |  | 0.096191 |  | 0.022235 |  | 0.011537 |  | 0.066983 |  | 0.91846 |  | 0.83512
T-test | 4.6691e-07 |  | 8.0967e-06 |  | 5.9253e-08 |  | 1.7123e-08 |  | 7.4756e-14 |  | 4.4893e-09 |  | 8.6974e-09 |  | 5.2639e-14 |  | 7.423e-17 |  | 7.4125e-17
Samples | Coated PET fibres (LN021/117/7) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/7 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2670.5381 | 2969.9845 | 2908.0942 | 1550.184 | 2426.572 | 2791.3984 | 699.6462 | 747.8233 | 1052.0071 | 697.2084 | 644.4613 | 639.8262
Image 2 | 2538.2256 | 2206.1071 | 2426.5574 | 2732.1684 | 2462.8707 | 2989.8978 | 647.8825 | 857.5156 | 494.3263 | 1261.1208 | 727.7484 | 625.3321
Image 3 | 2661.4379 | 2936.4061 | 1892.6389 | 2447.9203 | 2501.7456 | 2506.7508 | 669.1921 | 724.5108 | 693.9711 | 1339.6682 | 653.6579 | 664.9418
Image 4 | 2344.0255 | 2682.9015 | 2239.6627 | 2503.9035 | 2944.1804 | 2644.7199 | 781.5203 | 834.7282 | 995.7059 | 1614.0366 | 786.2552 | 515.9295
average | 2553.5568 | 2698.8498 | 2366.7383 | 2308.5441 | 2583.8422 | 2733.1917 | 699.5603 | 791.1445 | 809.0026 | 1228.0085 | 703.0307 | 611.5074
SD | 152.1637 | 352.6078 | 423.3056 | 520.3071 | 242.1787 | 206.8726 | 58.6239 | 64.8625 | 262.1467 | 384.8509 | 66.8464 | 65.7862
CV | 0.059589 | 0.13065 | 0.17886 | 0.22538 | 0.093728 | 0.075689 | 0.083801 | 0.081986 | 0.32404 | 0.31339 | 0.095083 | 0.10758
F-test | 0.0038253 |  | 0.00013665 |  | 0.01033 |  |  |  | 0.00049127
T-test | 4.1367e-08 |  | 1.3713e-05 |  | 5.3365e-09 |  |  |  | 0.03084
Samples | Coated PMP fibres (LN021/117/8) | [11 empty] | Coated PMP fibres (LN021/117/8) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740128
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | SU
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 3095.7839 | 3167.4607 | 2130.2091 | 2747.9731 | 3156.2471 | 2072.8191 | 3092.4681 | 3225.6741 | 2713.2792 | 3005.8853 | 3510.9512 | 2635.1009 | 2799.9783 | 3260.1446 | 2815.7726 | 2983.1482 | 3063.0943 | 3124.4226 | 2856.8669 | 2830.7196 | 641.1342 | 614.098
Image 2 | 2912.5104 | 1943.0745 | 2493.5883 | 2538.7546 | 2085.8124 | 2723.5441 | 2761.3994 | 2696.4247 | 2598.9237 | 2790.6016 | 2838.6978 | 2204.4112 | 2786.1773 | 2488.959 | 2796.7442 | 2967.6468 | 2747.3045 | 2804.6165 | 2648.0512 | 3062.6978 | 683.2529 | 542.851
Image 3 | 3047.8169 | 2444.1377 | 2508.3127 | 3234.1741 | 2686.5986 | 2685.6692 | 2560.5123 | 2602.3134 | 2825.3077 | 2826.4378 | 3118.7363 | 3069.2065 | 2355.8649 | 3037.2888 | 2984.8281 | 2507.7189 | 2708.4978 | 2607.8505 | 3029.7321 | 2958.8099 | 780.1471 | 739.8454
Image 4 | 2600.7825 | 2156.9137 | 2703.8491 | 1915.8539 | 2570.8314 | 2814.7198 | 2418.8477 | 2808.8248 | 2743.9601 | 2458.523 | 2615.9937 | 2650.4803 | 2909.1341 | 2615.8681 | 2970.047 | 2969.4372 | 2995.5461 | 3055.3027 | 2450.4056 | 2659.4819 | 633.1879 | 854.5569
average | 2914.2234 | 2427.8966 | 2458.9898 | 2609.1889 | 2624.8724 | 2574.1881 | 2708.3069 | 2833.3093 | 2720.3677 | 2770.3619 | 3021.0948 | 2639.7997 | 2712.7887 | 2850.5651 | 2891.848 | 2856.9878 | 2878.6107 | 2898.0481 | 2746.264 | 2877.9273 | 684.4306 | 687.8378
SD | 222.9047 | 534.0734 | 239.2228 | 546.3643 | 439.5738 | 338.6058 | 292.135 | 274.8613 | 93.7513 | 228.2316 | 385.9485 | 353.1239 | 244.2226 | 359.8438 | 99.3186 | 232.9489 | 176.9066 | 237.2892 | 251.5055 | 173.8094 | 67.4869 | 137.7906
CV | 0.076489 | 0.21997 | 0.097285 | 0.2094 | 0.16746 | 0.13154 | 0.10787 | 0.097011 | 0.034463 | 0.082383 | 0.12775 | 0.13377 | 0.090026 | 0.12624 | 0.034344 | 0.081537 | 0.061456 | 0.081879 | 0.091581 | 0.060394 | 0.098603 | 0.20032
F-test | 0.00069046 |  | 0.0017217 |  | 0.0030378 |  | 0.017877 |  | 0.22081 |  | 0.0017248 |  | 0.011127 |  | 0.20418 |  | 0.10364 |  | 0.066886
T-test | 3.1738e-06 |  | 1.5649e-06 |  | 5.1309e-07 |  | 8.9859e-09 |  | 3.6009e-14 |  | 5.041e-07 |  | 2.4093e-08 |  | 1.8796e-14 |  | 8.4363e-14 |  | 3.6693e-13
Samples | Coated PET fibres (LN021/117/8) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/8 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2239.3163 | 3517.918 | 1923.3508 | 1662.1972 | 2082.0846 | 2384.3404 | 699.6462 | 747.8233 | 1153.5803 | 809.1167 | 644.4613 | 639.8262
Image 2 | 2321.9232 | 2640.1952 | 2052.3677 | 1365.7382 | 2751.2392 | 2239.3069 | 647.8825 | 857.5156 | 1023.7608 | 901.6169 | 727.7484 | 625.3321
Image 3 | 2789.219 | 2250.0076 | 1723.6946 | 1797.9531 | 2772.5923 | 2013.4861 | 669.1921 | 724.5108 | 3130.3507 | 635.6919 | 653.6579 | 664.9418
Image 4 | 2545.7434 | 2544.8501 | 1144.2875 | 2375.0794 | 2449.3175 | 2530.5654 | 781.5203 | 834.7282 | 2825.1579 | 927.9642 | 786.2552 | 515.9295
average | 2474.0505 | 2738.2427 | 1710.9252 | 1800.242 | 2513.8084 | 2291.9247 | 699.5603 | 791.1445 | 2033.2124 | 818.5974 | 703.0307 | 611.5074
SD | 246.7886 | 545.6742 | 401.2268 | 423.593 | 323.4641 | 220.444 | 58.6239 | 64.8625 | 1099.0352 | 132.1585 | 66.8464 | 65.7862
CV | 0.099751 | 0.19928 | 0.23451 | 0.2353 | 0.12867 | 0.096183 | 0.083801 | 0.081986 | 0.54054 | 0.16145 | 0.095083 | 0.10758
F-test | 0.00019704 |  | 0.00033343 |  | 0.0024641 |  |  |  | 8.0062e-07
T-test | 2.9868e-06 |  | 0.0001149 |  | 2.3146e-07 |  |  |  | 0.06074
Samples | Coated PMP fibres (LN021/117/9) | [11 empty] | Coated PMP fibres (LN021/117/9) continued | [7 empty] | Uncoated PMP LN021/11/29 |  | S/N | 0740129
| layer 1 |  | layer 11 |  | layer 21 |  | layer 31 |  | layer 41 |  | layer 51 |  | layer 61 |  | layer 71 |  | layer 81 |  | layer 91 |  |  |  | Cond | ETO,TC
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B | side A | side B
Image 1 | 2402.0791 | 2430.7695 | 2179.5512 | 2602.575 | 2153.894 | 2361.8144 | 2723.5302 | 2312.8197 | 2659.9766 | 2555.2675 | 2450.9157 | 2431.7787 | 2790.2032 | 2571.534 | 2596.5436 | 2949.0614 | 3023.6776 | 3196.0131 | 3048.9249 | 2779.4708 | 641.1342 | 614.098
Image 2 | 2498.7252 | 2437.6305 | 2298.9324 | 1849.8054 | 2186.873 | 2344.4161 | 2442.8526 | 2222.2723 | 2586.3926 | 2376.3851 | 2720.0649 | 2627.6369 | 2810.5419 | 2526.649 | 2729.7401 | 2771.6964 | 2736.9762 | 3080.2874 | 3072.434 | 3283.3514 | 683.2529 | 542.851
Image 3 | 2618.2932 | 3115.4959 | 2426.7429 | 2249.3206 | 1516.764 | 2348.8483 | 2503.8983 | 2146.6767 | 2601.8216 | 2437.0937 | 2824.923 | 2344.6347 | 2580.8104 | 2309.374 | 2969.6193 | 2884.5608 | 2985.4714 | 2600.3209 | 3025.0942 | 3214.2122 | 780.1471 | 739.8454
Image 4 | 2399.5229 | 2595.5032 | 2550.7701 | 2063.2099 | 2228.1478 | 2332.0186 | 2311.2363 | 2167.7044 | 2327.1814 | 2763.7143 | 2367.2747 | 2612.6393 | 2380.5653 | 2680.023 | 3005.6651 | 2951.9016 | 2961.8915 | 2683.7408 | 2983.7179 | 2977.6618 | 633.1879 | 854.5569
average | 2479.6551 | 2644.8498 | 2363.9992 | 2191.2277 | 2021.4197 | 2346.7743 | 2495.3793 | 2212.3683 | 2543.843 | 2533.1151 | 2590.7946 | 2504.1724 | 2640.5302 | 2521.895 | 2825.392 | 2889.3051 | 2927.0042 | 2890.0905 | 3032.5428 | 3063.6741 | 684.4306 | 687.8378
SD | 103.3173 | 322.8586 | 160.2861 | 319.1338 | 337.8057 | 12.299 | 172.0398 | 74.1593 | 147.8752 | 170.7348 | 216.8365 | 138.6863 | 202.0348 | 155.6238 | 195.6364 | 84.3474 | 129.2181 | 292.2956 | 37.8552 | 230.2821 | 67.4869 | 137.7906
CV | 0.041666 | 0.12207 | 0.067803 | 0.14564 | 0.16711 | 0.0052408 | 0.068943 | 0.03352 | 0.058131 | 0.067401 | 0.083695 | 0.055382 | 0.076513 | 0.061709 | 0.069242 | 0.029193 | 0.044147 | 0.10114 | 0.012483 | 0.075165 | 0.098603 | 0.20032
F-test | 0.036014 |  | 0.027242 |  | 0.014372 |  | 0.10186 |  | 0.32839 |  | 0.1673 |  | 0.15197 |  | 0.36652 |  | 0.070168 |  | 0.28447
T-test | 4.0623e-09 |  | 3.6408e-08 |  | 2.4393e-07 |  | 3.9592e-12 |  | 5.7974e-14 |  | 2.8107e-13 |  | 2.752e-13 |  | 4.8679e-15 |  | 1.7929e-13 |  | 2.89e-15
Samples | Coated PET fibres (LN021/117/9) | [5 empty] | Uncoated PET LN021/11/29 |  | Coated H.P. |  | Uncoated H.P.
| layer 1 |  | layer 10 |  | layer 20 |  |  |  | LN021/117/9 |  | LN021/11/29
Surfaces | side A | side B | side A | side B | side A | side B | side A | side B | part A | part B | part A | part B
Image 1 | 2548.3963 | 2423.8289 | 2310.2117 | 2141.8531 | 3008.9988 | 2779.9615 | 699.6462 | 747.8233 | 1397.8987 | 1314.8711 | 644.4613 | 639.8262
Image 2 | 2097.6937 | 2680.7638 | 2100.1077 | 2383.3186 | 2617.1659 | 3069.8805 | 647.8825 | 857.5156 | 670.1315 | 1206.7186 | 727.7484 | 625.3321
Image 3 | 1981.3369 | 2355.3207 | 2239.9608 | 2432.4814 | 2699.3696 | 3108.864 | 669.1921 | 724.5108 | 2948.6673 | 657.7086 | 653.6579 | 664.9418
Image 4 | 2160.0307 | 2189.8829 | 1994.6713 | 2197.5245 | 2314.4712 | 2112.6929 | 781.5203 | 834.7282 | 2989.1352 | 1870.4492 | 786.2552 | 515.9295
average | 2196.8644 | 2412.4491 | 2161.2379 | 2288.7944 | 2660.0014 | 2767.8497 | 699.5603 | 791.1445 | 2001.4582 | 1262.4369 | 703.0307 | 611.5074
SD | 245.7765 | 204.0602 | 141.2702 | 140.8341 | 285.5258 | 460.7568 | 58.6239 | 64.8625 | 1156.0602 | 497.0678 | 66.8464 | 65.7862
CV | 0.11188 | 0.084586 | 0.065365 | 0.061532 | 0.10734 | 0.16647 | 0.083801 | 0.081986 | 0.57761 | 0.39374 | 0.095083 | 0.10758
F-test | 0.0069223 |  | 0.097582 |  | 0.00052243 |  |  |  | 1.2379e-06
T-test | 6.4839e-08 |  | 4.3638e-13 |  | 5.8316e-07 |  |  |  | 0.019423

## ANOVA
0740093 | 0740128 | Ttest
2378.4816 | 3095.7839 | 2.9457e-08
2341.7996 | 2912.5104
1869.9868 | 3047.8169
1260.284 | 2600.7825
2254.315 | 3167.4607
1627.5081 | 1943.0745
2227.8814 | 2444.1377
1786.6794 | 2156.9137
3265.7236 | 2130.2091
3025.1484 | 2493.5883
2907.7031 | 2508.3127
3227.0805 | 2703.8491
2562.7405 | 2747.9731
2774.7128 | 2538.7546
1222.9905 | 3234.1741
930.9537 | 1915.8539
2413.026 | 3156.2471
2315.0452 | 2085.8124
1968.8736 | 2686.5986
2073.3394 | 2570.8314
1696.616 | 2072.8191
1797.7282 | 2723.5441
1333.5361 | 2685.6692
1252.4967 | 2814.7198
2178.5054 | 3092.4681
1678.4096 | 2761.3994
1601.4655 | 2560.5123
1402.8272 | 2418.8477
2424.2884 | 3225.6741
1725.3871 | 2696.4247
1910.9667 | 2602.3134
2219.6376 | 2808.8248
2418.9599 | 2713.2792
2038.941 | 2598.9237
1786.8704 | 2825.3077
2592.8028 | 2743.9601
1598.537 | 3005.8853
2435.637 | 2790.6016
1822.8851 | 2826.4378
2050.4665 | 2458.523
2852.2896 | 3510.9512
2694.0107 | 2838.6978
3045.5644 | 3118.7363
2587.9352 | 2615.9937
1935.9269 | 2635.1009
2171.5758 | 2204.4112
2920.9688 | 3069.2065
1892.7111 | 2650.4803
1892.7323 | 2799.9783
2643.1281 | 2786.1773
2183.2422 | 2355.8649
2324.2728 | 2909.1341
2334.7448 | 3260.1446
2685.5554 | 2488.959
1629.778 | 3037.2888
1814.142 | 2615.8681
2053.7296 | 2815.7726
2664.7203 | 2796.7442
2445.6591 | 2984.8281
2776.003 | 2970.047
2390.2251 | 2983.1482
2307.5829 | 2967.6468
1964.8862 | 2507.7189
2172.4622 | 2969.4372
2912.6675 | 3063.0943
2614.8823 | 2747.3045
3180.8275 | 2708.4978
2792.6245 | 2995.5461
2417.3037 | 3124.4226
3263.8718 | 2804.6165
2865.0431 | 2607.8505
2760.3326 | 3055.3027
3258.5025 | 2856.8669
3146.5223 | 2648.0512
3319.7046 | 3029.7321
2704.0868 | 2450.4056
2989.9897 | 2830.7196
2788.713 | 3062.6978
3168.7609 | 2958.8099
3152.3975 | 2659.4819

## Serial No
Group | S/N | ref | PC Coated | EtO | Temp Cycle | 2 Yrs Simulated Age | 14 Day Sim Use
1 | ECMO (0740099) | LN021/117/3 | X
1 | ECMO (0740127) | LN021/117/7 | X
2 | ECMO (0740093) | LN021/117/2 | X |  |  |  | X
2 | ECMO (0740128) | LN021/117/8 | X |  |  |  | X
3 | ECMO (0740124) | LN021/117/6 | X | X | X
3 | ECMO (0740129) | LN021/117/9 | X | X | X
4 | ECMO (0740104) | LN021/117/4 | X | X | X |  | X
4 | ECMO (0740106) | LN021/117/5 | X | X | X |  | X
| Group # | PC Coated | EtO | Temp Cycle | 2 Yrs Simulated Age | 14 Day Sim Use | Number of device | S/N 1 | S/N 2 | Test
| 1 | X |  |  |  |  | 2 | 0740099 | 0740127 | Coverage
| 2 | X |  |  |  | X | 2 | 0740128 | 0740093
| 3 | X | X | X |  |  | 2 | 0740124 | 0740129
| 4 | X | X | X |  | X | 2 | 0740104 | 0740106
| 5 | X | X | X | X |  | 2
| 6 | X | X | X | X | X | 2
- Total 12 devices to be tested for coverage
- Request quote for testing coating coverage on Groups 1-4 from Vertellus (N=8)
- Groups 5 & 6 to be tested at later date (N=4, due back from Simulated Age 8/7/17)